WHOCALIB

WHO Thromboplastin Calibration Program.

Revised Version 3.1 (0902NC)

This program was prepared by:

The European Concerted Action on Anticoagulation (ECAA) Central Facility,

Stopford Building, The University of Manchester, Manchester M13 9PT, UK.

WHOCALIB, the WHO thromboplastin calibration program, may be freely used and distributed but copies should include this notice and all other files and documents associated with the original distribution package without alteration.  The present version, prepared by N Chauhan, has been based on a previous version by D Houghton.

This document has been edited by the ECAA Co-ordinator, L Poller.  Gratitude is expressed to DJ Finney, M Keown, V Chantarungkul, AMHP van den Besselaar, A Tripodi and J Jespersen for their helpful comments. Funding has been provided by EC Biomed Programme (Grant number PL931349) EC SMT Programme (Grant number 4-CT98-2269) and the Manchester Thrombosis Research Foundation.

Disclaimer

Please use this software at your own risk.  The ECAA disclaim any liability for direct, special, indirect, incidental or consequential damages believed to have been caused by the use of this program or any files derived from its use.

Minimum System Requirements

- 486-based personal computer

- Microsoft Windows 95

- 4 MB application RAM

- 4 MB hard disk space, plus 2 MB additional temporary disk space available during installation

1.  Installation Instructions

For installation from floppy disks:

- Insert Disk 1 into the floppy disk drive and ensure that all other applications on the system have been closed before you commence installation.

- Click the “Start” button and select “Run”. On the command line type “a:\setup” and then click “OK”.

- Insert Disk 2 when prompted, then click “OK”.

For installation from the downloaded files:

First ensure that the files

- Setup.exe

- Setup.lst

- WHOCALIB 3.1.CAB

are saved in the folder “c:\temp”. If this folder does not exist then it can be created using Windows Explorer (see Windows Help for instructions on this).

Then click the “Start” button and select “Run”. On the command line type “c:\temp\setup” and then click “OK”.

The program will be installed in the folder WHOCALIB 3.1 (the program can be installed to a different folder by clicking “Change Directory”). To begin the installation click on the button shown below.

Click here to install

- When the window shown below is displayed, click on “Continue”



After successful installation the following window will appear. Click “OK”.

2.  Using WHOCALIB

To use  WHOCALIB, once installation has been completed, select “WHOCALIB 3.1” from the program menu.

# INCLUDE PICTURE HERE

2.1.  Inputting data for a calibration

Calibration data are entered in 3 columns. The first of these columns must contain the abnormal or normal tag, the second must contain the prothrombin times (PT) in seconds of the reference thromboplastin and the third column the PT of the local or working thromboplastin, which is to be calibrated. Data may be entered in any order, provided correct tag is used.

To enter a PT result (in the columns labelled y or x), first click on the destination cell. Then click on the data input box (the white box on the top left hand of the Data Editor window). Enter the data and finally press “ENTER” on the keyboard.

Data Input Box

Destination cell

2.2.  Deleting a row of data

To delete a row of data, first select the row to be deleted by clicking on the grey area on the left hand side of the row.

Click here to select the third row

Then select the “Delete Row” option from the “Edit” drop-down menu.

2.3.  Opening an existing data set

An example test file, named “1999 who guidelines data-set.txt”, is provided.  This file contains the data analysed previously in the WHO Guidelines (1999). The first 10 lines of this file are given below. The first and second lines give the names of the reference and working thromboplastin respectively. The third line gives the ISI of the reference thromboplastin (here it is 1.0). The fourth line allows the standard error of the ISI of the reference thromboplastin to be entered (here it is set to zero). Reading from left to right, each of the remaining lines gives the abnormal or normal tag (1 and 0 represent normal and abnormal samples respectively), the PT with the reference thromboplastin and the PT with the working thromboplastin. The three values are separated by commas.

"RBT/90"

"secondary"

1

0

1,16.9,12.2

1,17.5,12.4

0,62.6,33.4

0,59.9,34.7

0,58.9,32.7

0,51.8,28.2

To open a file of data, saved in the format shown above, select the “Open” option from the “File” drop-down menu.

Then select the text file which you wish to open and click on “Open”.

The data will be displayed in the Data Editor window.

Residual SD limit of 3

2.4.  Calibration settings

Before performing the calibration, the calibration settings may need to be changed.

The box labelled “Calibration identifier” gives a name to identify the data-set. If an existing data-set has been opened then this is the name of the file containing the data-set. Otherwise, the identifier is set to “dataset1.txt”. The identifier can be changed by the user.

The boxes labelled “Reference Preparation (y)” and “Working Preparation (x)” give the names of the reference and working thromboplastins respectively.

The box labelled “ISI” gives the ISI of the reference thromboplastin. The standard error of this ISI is given in the box labelled “ISI Standard Error”. If the standard error is not known or the user does not wish to enter it, then the default value of zero will be used.

The boxes labelled “Lower INR” and “Upper INR” give the range of valid INR with the reference thromboplastin. The default values of 1.5 and 4.5 give the range of valid INR recommended by WHO Guidelines (1999). Abnormal samples with INR outside the set range will be excluded from the calibration.

According to WHO Guidelines (1999), the logarithm (log) of the PT of normal and abnormal samples obtained with the reference thromboplastin are plotted on the vertical axis against the log PT obtained with the working thromboplastin on the horizontal axis. An initial calibration line is derived by orthogonal regression analysis. The standard deviation of the log PT, termed “residual SD” in the WHO Guidelines (1999), is then calculated and samples are excluded if the perpendicular distance from the initial calibration line is greater than 3 times the residual SD. The default value of 3 given in the box labelled “Limits about calibration line” sets the residual SD limit to 3, which is the limit recommended by WHO Guidelines (1999). The residual SD limit can be changed by the user.

To perform the calibration, at least three abnormal samples and two normal samples are required after exclusions. The final calibration line and the mean normal prothrombin time (MNPT) (geometric mean of the normal PT) are then determined from the samples remaining after exclusions.

The box labelled “Confidence Intervals” allows the user to select either 90%, 95% or 99% confidence intervals.

The box labelled “Tick box for Tomenson’s corrected INR” can be ticked to display Tomenson corrected INR (see section 3).

The calibration is performed by pressing the “Calibrate” button.

2.5.  Results

Results are displayed in the window labelled “Calibration Results”. The calibration results for the example test file, “1999 who guidelines data-set.txt”, are given below. An interpretation of the results marked [A] to [H] is provided below the output.

WHO THROMBOPLASTIN CALIBRATION PROGRAM

Version 3.1 (0902NC)

__________________________________________________________

File Name: C:\Program Files\WHOCALIB 3.1\1999 who guidelines data-set.txt

Date: 17/09/02                                             Time: 11:36:22

Reference thromboplastin: RBT/90

Working thromboplastin: secondary

20  normal observations and

[A]

60  abnormal observations in original data-set

__________________________________________________________

Observations outside INR limits

(MNPT derived from all 20 normal observations

(lower limit=1.50; upper limit=4.50)

Plasma     PT           INR

[B]

27       83.4         4.71

43       87           4.91

51       94.3         5.32

__________________________________________________________

ABNORMAL+NORMAL CALIBRATION (n = 77)

Coefficients  :          y  =  -0.4751  +  1.3138 x

Standard error:               (0.0489)    (0.0151)

[C]

SD about line = 0.0330   CV of slope = 1.15

95% Confidence Interval for slope: (1.2841, 1.3444)

__________________________________________________________

ABNORMAL ONLY CALIBRATION (n = 57)

Coefficients  :          y  =  -0.8028  +  1.4083 x

Standard error:               (0.2084)    (0.0605)

[D]

SD about line = 0.0357   CV of slope = 4.30

95% Confidence Interval for slope: (1.2934,  1.5371 )

__________________________________________________________

TOMENSON'S CORRECTION

Vertical displacement, d, between abnormal only line

[E]

and mean of normal log PT = -0.0945 (exp(d) = 0.9098)

95% Confidence Interval for displacement: (-0.205,  0.016 )

Test for zero displacement: test statistic = -1.71

degrees of freedom =  57

0.090 < p-value < 0.100

ISI STATISTICS

ISI (RBT/90) = 1.0000

[F]

MNPT =  17.7174  ( 20 normals)

Standard error = 0

CV of ISI =  0

ISI (secondary) = 1.3138

MNPT = 12.7320 ( 20 normals)

[G]

Standard error = 0.0151

CV of ISI =  1.15

PROTHROMBIN TIMES AND INTERNATIONAL NORMALIZED RATIOS

Reference           Working             Tomenson's

Thromboplastin      Thromboplastin      corrected

Plasma       PT        INR       PT        INR       INR

1        16.90     0.95      12.20     0.95      0.86

[H]

2        17.50     0.99      12.40     0.97      0.88

3        62.60     3.53      33.40     3.55      3.54

4        59.90     3.38      34.70     3.73      3.73

5        58.90     3.32      32.70     3.45      3.43

[A]

The number of normal and abnormal samples in the original data-set

[B]

Abnormal samples which have been excluded because their INR with the reference thromboplastin are outside the range 1.5 to 4.5. The number of the sample, the PT and INR are provided.

None of the samples had a perpendicular distance from the initial calibration line greater than 3 times the residual SD. If the “Limits About Calibration Line” had been set to 2 then three samples would have been excluded because their perpendicular distance from the initial calibration line exceeded 2 times the residual SD. The following would then have been output.

Observations outside SD limits

(SD limit=2.00)

Plasma     PT

37       43.9

52       70.5

75       75.6

[C]

- The calibration line derived with both abnormal and normal samples;

- n is the total number of samples used to fit the calibration line;

- y is the log PT with the reference thromboplastin;

- x is the log PT with the working thromboplastin;

- -0.4751 and 1.3138 are the intercept and slope of the calibration line;

- the values given in parentheses immediately below the intercept and slope are the corresponding standard errors;

- the “SD about line” value is the residual SD (a measure of dispersion of the points about the calibration line);

- the ”CV of the slope” gives the coefficient of variation (%) of the slope of the calibration line

i.e. 100 x slope standard error/slope = 100 x 0.0151/1.3138 = 1.15;

- a confidence interval for the slope is also given.

[D]

The calibration line derived with abnormal samples only (i.e. without the normal samples).

[E]

- The section marked “Tomenson’s correction” gives the displacement (labelled “d”). This is the vertical distance of the mean log PT of the normal samples used in the calibration from the abnormal only calibration line. The exponential of the displacement, “exp(d)”, is also provided.

- A confidence interval for the displacement is given. A test for zero displacement (Tomenson, 1984) is performed and the test statistic given in the output. The degrees of freedom and a range for the p-value are also given.

[F]

- “ISI (RBT/90)” gives the ISI of the reference thromboplastin.

- “MNPT” ; gives the MNPT with the reference thromboplastin. This is calculated from the normal samples used to derive the abnormal+normal calibration line. The total number of normal samples used to calculate the MNPT is given in parentheses.

- The standard error of the ISI of the reference thromboplastin had been set to zero hence the “CV of ISI” is also zero (see section 3).

[G]

- “ISI (secondary)” gives the ISI of the working thromboplastin;

- “MNPT” ; gives the MNPT with the working thromboplastin. This is calculated from the normal samples used to derive the abnormal+normal calibration line. The number of normal samples used to calculate the MNPT is given in parentheses.

- The standard error and CV of the ISI of the working thromboplastin are calculated using the standard error of the ISI of the reference thromboplastin.

[H]

- The PT and INR with both thromboplastins are provided for all the abnormal and normal samples. Note that the output given shows the results for only the first 5 samples.

- By ticking the box labelled “Tick box for Tomenson’s corrected INR” on the Data Editor window, Tomenson’s corrected INR are also displayed (see section 3 for further details).

2.6.  Saving Data

Data entered in the Data Editor window (see section 2.1) can be saved as a text file by selecting the “Save Data As…” option from the “File” drop-down menu.

2.7.  Saving Results

The contents of the Calibration Results window can be saved as a text file by selecting the “Save Results As…” option from the “File” drop-down menu.

2.8.  Plotting the PT data and the calibration lines

After pressing the “Calibrate” button on the Data Editor window (see section 2.1), a plot of the PT data with the abnormal+normal and abnormal only calibration lines can be produced by selecting the “Produce Plot” option from the “Plot” drop-down menu.

The default plot, corresponding to the example test file, “1999 who guidelines data-set.txt”, is shown below.

The plot can be copied for pasting elsewhere (for example, into a MS Word document) by selecting the “Copy Plot” option from the “Plot” drop-down menu.

The plot can be saved as a “bitmap” file by selecting the “Save Plot” option from the “Plot” drop-down menu.

Some features of the plot can be changed by selecting the “Plot Options” option from the “Plot” drop-down menu.

The plot title and axis labels can be changed. Also, the abnormal only line and/or the legend can be excluded from the plot. To save the changed plot settings press the “Apply” button. To revert to the previous plot settings press the “Cancel” button. After changing the plot settings, select “Produce Plot” from the “Plot” drop-down menu.

2.9.  Switching between windows.

The Data Editor, Results and Plot windows can be viewed by selecting the corresponding option from the “View” drop-down menu.

2.10.  Printing calibration results and plots

The contents of the Calibration Results window and/or the plot can be printed by firstly selecting the “Print” option from the “File” drop-down menu.

Then tick the “Results” and “Plot” boxes to print the calibration results and plot, respectively. Finally click the “OK” button.

2.11.  Resetting WHOCALIB 3.1

To reset WHOCALIB 3.1 (i.e. to remove the contents of the Data Editor, Calibration Results and Plot windows) select the “New” option from the “File” drop-down menu. It is recommended that this be done before a new data-set is entered.

2.12.  Quitting WHOCALIB 3.1

To exit WHOCALIB select “Exit” from the “File” drop-down menu. Data in the Data Editor window, which has not been saved, will be lost.

3.  Calculations

The calculations for calibration of the working thromboplastin are described in this section. Many of the calculations have been described previously in the WHO Guidelines (1999). For additional calculations, relevant references have been provided.

Let xi and yi represent the natural logarithm of the PT of the ith sample with the working and reference thromboplastins respectively. Let the total number of abnormal and normal samples be represented by nA and nN respectively and let n = nA + nN be the total number of samples. Assume that the first nN observations are for the normal samples and the remaining nA observations are for the abnormal samples.

The first step of the calibration requires the calculation of the MNPT with the reference thromboplastin (denoted by My). This is the geometric mean PT of the normal samples, calculated as follows:

My = exp1nNi=1nNyi

The INR of the ith abnormal sample with the reference thromboplastin is then calculated as

INRi = exp(yi)MyISIref,

where ISIref is the ISI of the reference thromboplastin.

The ith abnormal sample is then excluded if INRi is outside the valid INR range.

Now let nA represent the number of abnormal samples after exclusion.

The yi are plotted on the vertical axis against the xi on the horizontal axis. A straight line, of the form

y = a + bx

is fitted using orthogonal regression analysis. The estimation of the intercept (a) and slope (b) is now described.

First calculate the arithmetic means of xi and yi:

x = 1ni=1nxi	and	y = 1ni=1nyi.

Then calculate the following:

Qxx = i=1nxi−x2, 		Qyy = i=1nyi−y2	and	Qxy = i=1nxi−xyi−y.

Now let:

E = Qyy−Qxx2+4Qxy2.

The slope and intercept can then be estimated as follows:

b = Qyy−Qxx+E2Qxy

a = y - bx.

The “SD about the line” (called the “residual standard deviation in the WHO Guidelines, 1999) is calculated as V, where

V = Qyy−bQxy(n−2).

The perpendicular distance (wi) of the point (xi, yi) from the fitted line y = a + bx is then calculated as follows:

wi = (yi−y)−b(xi−x)(1+b2).

For a residual SD limit, denoted by F, any samples (abnormal or normal) with wi exceeding FV are then excluded. The default value of F=3, results in samples being excluded if the perpendicular distance from the calibration line is greater than 3 times the residual SD.

Now let nA and nN represent the number of abnormal and normal samples after exclusions.

The calibration line is then re-fitted (called the “final” calibration line) using the remaining samples (i.e. x, y, Qxx, Qyy, Qxy, E, a, b and V are re-calculated using the remaining samples). The standard error of the slope [denoted se(b)] and the standard error of the intercept [denoted se(a)] are then calculated as follows:

se(b) = [(1+b2)Qxy+nbV]bVQxy2

se(a) = se(b)x2+(1+b2)Vn          (van der Velde, 1984).

The CV of the slope is given by

100 x se(b)b.

The confidence interval for the calibration slope is calculated as:

tan(g - h)      and    tan(g + h),

where

g = tan−1(b), 	h = 12tan−1G(1−G)

and	G = 4t2QxxQyy−QxyE(n−2).

For a 100(1-α)% confidence interval, t  is the upper α2 percentage point of the 
t-distribution with n-2 degrees of freedom. For the example test file, “1999 who guidelines data-set.txt”, n=77 and t = 1.992 for a 95% confidence interval (i.e. when α = 0.05).

The calibration is also performed using only the abnormal samples used to fit the final calibration line. The intercept and slope of this line (referred to as the “abnormal only line”) are denoted by aA and bA respectively.

The calculation of the results marked “Tomenson’s correction” is now outlined (see Tomenson (1984) for further details). Let xN and yN be the arithmetic means of the log PT of the normal samples used to fit the final calibration line:

xN = 1nNi=1nNxi 	and		yN = 1nNi=1nNyi.

The displacement (d) is the vertical distance of the point (xN, yN) from the abnormal only calibration line:

d = aA + bAxN - yN.

(This corrects an error in Tomenson (1984), where the displacement was defined as –d).

The confidence interval for the displacement is calculated as:

d ±tS1+S2,

where

S1 = 1nN(nN−1)Qyy(N)−2bAQxy(N)+bA2Qxx(N),

S2 = se(aA)2+xNse(bA)2xN−2xA,

Qxx(N) = i=1nNxi−xN2,     Qyy(N) = i=1nNyi−yN2,       Qxy(N) = i=1nNxi−xNyi−yN

and 	xA = 1nAi=nN+1nxi  is the arithmetic mean of the log PT, with the working thromboplastin, of the abnormal samples used to fit the final calibration line.

For a 100(1-α)% confidence interval for the displacement, t  is the upper α2 percentage point of the t-distribution with f degrees of freedom, where f is the integer part of

S1+S22S12(nN+1)+S22nA−2.

For the test for zero displacement at the 100α% significance level, the test statistic
	dS1+S2

is compared to a t-distribution with f degrees of freedom (Welch, 1947).

(Biometrika Table 11, referred to by Tomenson (1984), gives approximately the same results as the 
t-distribution).

The MNPTs reported in the “ISI Statistics” section of the output are the geometric means of the log PT of the normal samples, which were used to fit the final calibration line.

The ISI of the working thromboplastin is

ISIwork = b(ISIref).

The standard error of ISIworkis calculated as:

(ISIref)se(b)2+se(ISIref)b2,

where

se(ISIref) is the standard error of the ISI of the reference thromboplastin.

The calculation of the INR is now described. The INR with the reference thromboplastin is

exp(yi)MyISIref

and the INR with the working thromboplastin is

exp(xi)MxISIwork,

where

My = exp(yN)     and     Mx = exp(xN)

are the geometric mean PT of the normal samples.

If a significant displacement is observed, Tomenson (1984) recommends that a corrected INR (referred to as Tomenson’s corrected INR) be calculated as follows:

exp(d)exp(xi)MxbAISIref.

4.  References

Tomenson JA (1984). A Statistician's independent evaluation. In:

Thromboplastin calibration and oral anticoagulant control. van den Besselaar AMHP, Gralnick HR, Lewis SM (eds.). Martinus Nijhoff Publishers, Boston; pp 87-108.

van der Velde EA (1984). Orthogonal regression equation. In:

Thromboplastin calibration and oral anticoagulant control. van den Besselaar AMHP, Gralnick HR, Lewis SM (eds.). Martinus Nijhoff Publishers, Boston; pp 25-39.

Welch BL (1947). The generalisation of ‘Students’ problem when several different population variances are involved. Biometrika; 34: 28-35.

WHO Guidelines (1999). WHO guidelines for thromboplastins and plasma used to control oral anticoagulant therapy. Annex 3. WHO Technical Reports; 889: 64-93.